## Slide 1
Label-free three-dimensional cellular viability assay using optical coherence tomography
Cell therapy, boosted by recent advances in culture technology, has emerged as a promising approach for repairing damaged tissues and curing diseases. Viability measurement is a critical quality control step in cell therapy manufacturing to ensure the used cells are viable and functional for therapeutic purposes. However, traditional gold-standard methods, such as trypan blue and fluorescent staining, require labeling and pose contamination risks. Therefore, a label-free and viability-sensitive detection method is desired by pharmaceutical and biomedical researchers.
Optical coherence tomography (OCT) has gained attention in biological imaging due to its non-invasive, high-resolution, and three-dimensional (3D) imaging capabilities. Recent studies have demonstrated its potential for assessing viability of tissue and organoids, suggesting its potential for cell-based applications. In this study, we propose a label-free approach to estimate cellular viability using OCT. Unlike traditional methods, this approach is purely optical-based detection through the vial, hence has minimal contamination risk.
Robustness to aggregation:
Aggregated cells were observed in both viable and non-viable samples, posing challenges for effective mixing in NC measurements. NC relies on subsampling from monodisperse cell suspensions for accurate viability assay and therefore is inherently sensitive to such errors. Imperfect mixing with residual aggregation can result in unrepresentative subsamples, likely contributing to the low linearity and high variance seen in NC results.
In contrast, OCT is label-free and capable of acquiring 3D images, making it potentially high-throughput and suitable for assessing a larger and more representative portion of the cell population. This makes OCT potentially robust to aggregation, less prone to sampling variability, and better measurement consistency.
Limitations and future work：
While OCT can analyze dense cell sediments up to hundreds of microns deep, suspended cells may also be captured whose temporal speckle pattern is contributed by multiple phenomena hence not necessarily reflect their viability. Excluding suspending cells is necessary to assure the accuracy of viability assay by OCT. In this study, we manually selected depth ranges (40-100 μm) to exclude suspended cells. Centrifugation may help settle more cells and further improve the measurement accuracy.
This proof-of-concept study used only one cell type and a single viability-reduction method (heat shock). Future work will expand the approach to other cell types and viability-altering mechanisms to validate the versatility.
Background
Discussions
Methods
Logarithmic intensity variance (LIV) algorithm was used to quantify the temporal fluctuation magnitude of the OCT signal sequence. In this study, we employed a modified LIV (mLIV) by diving the standard LIV by the maximum signal intensity within a defined time window. This modification combines the temporal feature with signal intensity, mitigating the contrast loss due to the signal-to-noise ratio (SNR) discrepancy by the limited depth of focus. An empirical threshold of 0.056 on the mLIV was used to classify pixel-level viability. The impact of thresholding to the quantification is discussed later.
Viability was quantified as the proportion of pixels with mLIV value above the threshold within the segmented cell region, which was defined via signal-intensity-based masking. Suspended cells and low-SNR regions were manually excluded to ensure accuracy.
For visualization. each spatial pixel’s mLIV was encoded as hue, signal intensity as value, and saturation was set to 1 in HSV color space. In the resulting 3D images, yellow to green regions may represent viable cells, while red may indicate non-viable areas.
NucleoCounter (NC-200, ChemoMetec, Denmark), in short as NC, was used as a reference. It is a gold-standard fluorescence-based viability assay, using special cassettes to extract small volume of the sample and employing fluorescence staining to count the viability.
Experiment protocols
Sample preparations:
Jurkat cells (JKT-beta-del, JCRB0147, JCRB, Japan) were used to validate the proposed method. The cells were thawed at 37℃ and suspended in RPMI 1640 medium. Half of the sample was heat-treated at 55℃ for 150 minutes to reduce viability. After that, the treated and untreated cells were mixed to create five subsamples with varying viability levels made by different mixture ratios. All subsamples were transferred to AT-closed vials (Aseptic Technologies, Belgium) for measurements.
Measurement protocols:
NC measurements were performed first. The vials were vortexed gently before extracting a
Results
These findings suggest that OCT can provide sensitive and label-free measurements of cellular viability in comparison with NC. This approach demonstrates strong potential for manufacturing and preclinical workflows where it is preferable not to open the culture vessel.
(Left) The OCT light probed invertedly through the bottom of the vial to capture the sedimented cells for analysis. (Right) A representative 3D image of a viable (25%) and non-viable (75%) mixture of Jurkat cells. In this study, the lateral scanning area was 500 x 500 μm2 sampled by 256 x 256 locations. At each slow scan location 16 frames were repeated to capture sequential signals of the sample, whose time window was 0.3 s. The A-line rate was 20 kHz, frame rate was 50 FPS, and a volumetric acquisition took 70 s.
Wavelength:
84080 nm
Resolutions:
2.9 μm (axial)
5.8 μm (lateral)
Depth of focus:
360 μm (in air)
mLIV (dB)
Intensity (dB)
1
-5
10
0
Viable
Dead
En-face OCT images of heat-treated cells over time show a decreasing viable area (yellow to green) and an expanding non-viable region (red). The whole cell region shows red at 150-min.
X
Y
100 μm
150-min
0-min
60-min
120-min
(Left) Both OCT and NC indicate a consistent decrease in viability over time. OCT showed higher viability than NC in early time points (0-30 min), but lower viability at later stages (90-150 min). (Right) OCT demonstrated good linearity across the five mixture ratios, likely due to its purely imaging-based quantification. NC’s low linearity may have resulted from inconsistent vortexing, especially in the presence of aggregated cells.
NC
OCT
0
30
60
90
120
150
Time by heat shock (min)
Viability (%)
100
80
60
40
20
0
100
80
60
40
20
0
25
0
50
75
100
Viable cell mixture (%)
Viability (%)
subsample using the cassette for NC analysis. Each subsample was measured four times.
OCT measurements followed after a period of sedimentation for 7 minutes. Four OCT measurements were performed at similar locations.
To assess Sensitivity to viability, both NC and OCT measurements were conducted every 30 minutes during the heat treatment. To assess Linearity of measurement, subsamples with mixture ratio of 0%, 25%, 50%, 75%, and 100% of viable cells were measured by NC and OCT.
Viability (%)
100
80
60
40
20
0
0
30
60
90
120
150
Time by heat shock (min)
NC
OCT
OCT (+25%)
OCT (-25%)
Threshold optimization：
The viability threshold can directly impact the viability estimation by OCT. We tested several thresholds across all datasets, finding consistent decreasing tendency of viability. The threshold implemented in this study was selected to best match the readouts of NC. This degree of flexibility allows users to optimize the protocol based on their specific measurement and sample conditions, which is not available in conventional assays.
Coupler
Light source
Reference arm
Scanner
Interferometer
Sample
Camera
Grating
Spectrometer
OCT utilizes low-coherence interferometry to visualize 3D structure of samples by detecting back scattered light from within biological samples. The nature of coherence detection empowers it with high sensitivity to the refractive index change caused by the intracellular metabolic activities. By analyzing temporal fluctuations of OCT signal, we aimed to distinguish viable from non-viable cells.
Lida Zhu1*, Oliver Harriman1, Andy Atsushi Morosawa1
1)Santec LIS, Komaki, Japan
*lida.zhu@santec.com

## Slide 2
Time (min)
0
30
60
90
120
150
Viability (%)
0
20
40
60
80
100
0
25
50
75
100
0
20
40
60
80
100
(Left) The OCT light probed invert through the bottom of the vials to capture the sediment of cells for analysis. (Right) A representative 3D image of viable and non-viable mixture of Jurkat. The lateral scanning area was 500 x 500 μm2 sampled by 256 x 256 locations. At each slow scan location 16 frames were repeated to capture sequential signals of the sample, whose time window was 0.3 s. The Aline rate was 20 kHz, frame rate was 50 FPS, and a volumetric acquisition takes 70 s.
Wavelength:
84080 nm
Resolution:
2.9 μm (axial)
5.8 μm (lateral)
DOF:360 μm (air)
mLIV (dB)
Intensity (dB)
1
-5
10
0
Viable
Dead
(Left) Quantitative comparison between the OCT and NC viability measurements. Both OCT and NC show similar decreasing tendency of viability during the heat shock. OCT shows higher viability than NC at 0-min and 30-min, while lower from 90-min to 150-min. (Right) Five mixtures with various proportion of viable and non-viable cells were measured by OCT and NC. OCT shows a good linearity of viability to the mixture proportions, while NC shows less linearity than OCT possibly due to the non-perfect vortexing.
NC
OCT
0
30
60
90
120
150
Time of heat shock (min)
Viability (%)
100
80
60
40
20
0
100
80
60
40
20
0
25
0
50
75
100
Viable cell mixture (%)
Viability (%)
Viability (%)
100
80
60
40
20
0
0
30
60
90
120
150
Time (min)
NC
OCT
OCT (+25%)
OCT (-25%)
Dead
Viable
Coupler
Light source
Reference arm
Scanner
Interferometer
Sample
Camera
Grating
Spectrometer
En-face OCT images of heat-treated cells over time show a decreasing viable area (yellow to green) and an expanding non-viable region (red). The whole cell region shows red at 150-min.
X
Y
100 μm
150-min
0-min
60-min
120-min
mLIV (dB)
Intensity (dB)
1
-5
10
0
Viable
Dead

## Slide 3
mLIV (dB)
Intensity (dB)
1
-5
10
0
Viable
Dead
X
Y
Z
100 μm
Raster repeat:
(each Y) 16 repeats in 0.3 s, FPS=50
En-face sampling:
256 pixels (fast) X 256 pixels (slow)
Scanning area:
0.5 X 0.5 mm2
Voxel size (pixel separation):
2 X 2 X 4 μm (Z, X, Y)
Optical resolution (air):
Axial: 2.9 μm; lateral: 5.8 μm
Volume acquisition time: ~70 s
Wavelength:
84080 nm
Resolution:
2.9 μm (axial)
5.8 μm (lateral)
Depth of Focus:
360 μm (air)
Aline rate:
20 kHz
256 locations
X
16 repeats
X
Y
100 μm
256 locations
X
16 repeats
X
Y
100 μm
Dead
Viable
Viable
Non-viable
X
Z
OCT is a non-invasive imaging technique based on low coherence interferometry. It measures the interference of light backscattered from different layers within a sample to visualize its 3D structure. The nature of coherence detection empowers it with high sensitivity to the refractive index change caused by the intracellular metabolic activities. By monitoring the signal fluctuation, temporal features of OCT signal can be used to discriminate viable and non-viable cells (Fig. 1).
Logarithmic intensity variance (LIV) algorithm [??] was used to quantify the temporal fluctuation magnitude of the OCT signal sequence. In this study, we employed a modified LIV (mLIV):
where denotes the log scale signal intensity, is the spatial indices of an OCT volume, is the temporal index of a sequential OCT images, is the frame number of the image sequence. This modification combines the temporal feature with signal intensity, aiming to mitigate the contrast loss due to the signal-to-noise ratio (SNR) discrepancy by the limited depth of focus (DOF). An ad-hoc threshold of 0.056 was used to discriminate viable and non-viable (dead) cells. The impact of thresholding will be explored in Discussion ??.
Viability was defined as the fraction of number of pixels exhibiting mLIV value larger than an ad hoc threshold over the number of pixels within the whole cell region. The cell region was masked by applying an intensity threshold. Depth range was manually selected to ensure sufficient SNR and exclude suspending cells from analysis. Details about depth range and threshold selections will be discussed in Discussion ??.
Wavelength:
84080 nm
Resolution:
2.9 μm (axial)
5.8 μm (lateral)
DoF:360 μm (air)
Aline rate: 20 kHz
Viability (%)
100
80
60
40
20
0
0
30
60
90
120
150
Time (min)
NC
OCT
OCT (+25%)
OCT (-25%)
X
Y
100 μm
150-min
0-min
60-min
120-min
mLIV (dB)
Intensity (dB)
1
-5
10
0
Viable
Dead
NC
OCT
0
30
60
90
120
150
Time of heat shock (min)
Viability (%)
100
80
60
40
20
0
100
80
60
40
20
0
25
0
50
75
100
Viable cell mixture (%)
Viability (%)
Coupler
Light source
Reference arm
Scanner
Interferometer
Sample
Camera
Grating
Spectrometer